Fitting Keratoconics

Keratoconus represents a non-inflammatory thinning and cone-like protrusion of the cornea usually manifesting itself in youth or adolescence. The condition is usually bilateral, although the severity of involvement may be markedly asymmetrical.

Keratoconus usually results in considerable visual impairment, even with best spectacle correction, due to the development of a high degree of irregular astigmatism.  If the keratoconus is significant such that best spectacle acuity is reduced, the condition is best managed by fitting with rigid gas permeable contact lenses, as these provide a regular refracting surface over the cone.  Other contact lens options for the keratoconic patient include soft lenses, piggyback lenses and scleral lenses.

To successfully fit keratoconic patients, practitioners must have an appreciation of the different types of cone morphology.  Basically there are two types of cones.  Nipple cones are charactised by their small size and steep curvature.  The apical centre is most often located centrally or decentered slightly inferiorly.  Oval (or sagging) cones are larger, ellipsoid and commonly displaced inferotemporally.  Nipple cones are the easiest to fit.  Sagging cones are usually more difficult to fit due to the larger size of the cone and the often decentered location.

RGP fitting

Although many different fitting philosophies have been described for fitting RGP lenses to keratoconic patients, the two most common techniques used are the three-point touch and apical clearance methods.

The idea with the three-point touch technique is to distribute the weight of the lens between the cone and the more normal peripheral cornea, with the end result being light touch on the apex of the cone as well as the midperiphery such that a three-point touch fluorescein fitting pattern is observed.  It has been postulated that too much central touch can exacerbate or leads to corneal apical scarring, so care needs to be taken that the major lens bearing is on the normal peripheral `cornea with cone contact being limited to less than 3 mm.

With the apical clearance technique the lens is fitted to vault the central cornea, providing clearance over the apex.  The fluorescein pattern should show central clearance, intermediate bearing and peripheral clearance.  Fitting with apical clearance will generally lead to reduced acuity compared to the three point touch technique.

A rigid lens designed to fit a keratoconic cornea using the three point touch technique will generally have significantly different parameters from the RGP lens prescribed for a non-keratoconic cornea.  The back optic zone radius is usually steeper than normal due to the conical nature of the cornea.  The peripheral cornea is relatively normal so keratoconic eyes tend to be fitted with lenese have fairly normal back peripheral radii in combination with a steep BOZR, such that a greater axial edge lift is required compared to a non-keratoconic eye to achieve an acceptable degree (about 0.08 to 0.1mm) of edge clearance.  The back optic zone diameter is small than usual to allow adequate tear exchange under the cone and so minimize tear pooling and debris accumulation under the optic zone of the lens. The total lens diameter is usually smaller than normal so as to aid centration.  The required back vertex power is often highly negative due to the significant positive tear lens which results from fitting with a steeper BOZR.

Back optic zone radius

For an optimal three-point touch fluorescein pattern, there should be about 1-2 mm apical touch, an intermediate clearance zone, mid peripheral contact and then a normal band (approximately 0.5mm – 1 mm) of edge clearance.

If using keratometry readings to determine BOZR, a good starting point is a BOZR approximately half-way between the steepest and flattest corneal radius.  Work in big steps (~0.2 mm).  err on the steep (small) side of your final order.  If keratometry readings are off the scale, a BOZR of less than 6.8 mm is usually indicated.

While a keratoconic cornea is usually highly astigmatic, the irregularity of the astigmatism is s uch that generally it has asymmetry in any one meridian, and so centrally a toroidal back surface fits no better than a spherical one (especially as this meridional asymmetry is much greater centrally than it is peripherally). Consequently the use of a toric back optic zone should be avoided as it will not increase the area of alignment and it may also decrease acuity due to the toric BOZR spinning on the cone.

Lens diameter

The average lens diameter for a TGP lens fitted to a keratoconic eye is about 8.5 to 9 mm.  this is smaller than the average RGP lens size for a non-keratoconic eye and is mainly due to the fact that larger lenses will often be trapped in a low riding position as a result of a cone being located slightly inferiorly.

Generally, as the keratoconus progresses the required lens diameter decreases.  Note that it is not uncommon to have to prescribe different lens diameters for the two eyes if one eye is significantly more advanced than the other eye.

Back optic zone diameter

The back optic zone diameter is smaller than usual to allow adequate tear exchange around the cone and so minimising tear pooling and debris accumulation under the optic zone of the lens.  Note that the BOZD should not  be so small as to cause flare.

As the keratoconus progresses, the BOZD should be decreased.  A reasonable approximation is that the BOZD should equal BOZR although this does not hold as well in cases of very early and very advanced keratoconus.

Back peripheral radii

The aim is to achieve a nice, even band of clearance (approximately 0.5 mm – 1 mm) around the edge of the lens.  The peripheral cornea is relatively normal, hence keratoconic eyes tend to be fitted with lenses having fairly normal BPR in combination with a steep BOZR, such that a greater axial edge lift is required compared to a non-keratoconic eye.  AEL values in excess of 0.25 mm may be required with keratoconic eyes to achieve an acceptable degree (about 0.08 to 0.1 mm) of edge clearance. The greater the axial edge lift (that is, the steeper the BOZR), the greater the number of peripheral curves one should use to ensure a well balanced lens periphery.

Often there will be a significant degree of with the rule astigmatism in the peripheral cornea, such that the edge clearance along the vertical meridian will be greater than that along the horizontal (with inferior lens lift-off usually being noted).  As mentioned earlier, the irregularity of the astigmatism and the meridional asymmetry in the keratoconus generally precludes the use of a toric back optic zone.  Consequently, this problem is best solved by the use of the toric back peripheral radii in combination with a spherical optic zone. Anywhere between the 0.6 and 0.9 mm of toricity may be incorporated into the BPR and no form of lens stabilisation is needed.

Back Vertex Power

The required BVP is often highly negative due to the significant positive tear lens which results from fitting with a steeper BOZR.  The BVP should always be determined by refracting over a trial lens with an appropriate BOZR.

Be careful when calculating changes in BVP due to changes in the BOZR.  The normal rule of thumb is that we add -0.25 D to the BVP of the contact lens for every 0.05 mm decrease in BOZR.  Likewise, +0.25 D must be added to the BVP of the contact lens for every 0.05 mm increase in BOZR.  Unfortunately this approximation breaks down very quickly when working with the steeper (<7.00mm) BOZR. Consequently it is much safer to use the appropriate formula when modification in BVP is required as a result of a change in BOZR.